## Instructions
Data Entry and Analysis
The spreadsheets are designed to CLSI recommendations. The cells highlighted in yellow are cells that require input of information and data.
1. The first worksheet in the workbook is a Test Pool Calculator. Use this sheet to verify concentrations of interferent in the Assurance Interference Kits and also to calculate the required amount of Interference Material to add to the Base Pool to create the Test Pools. The section on the right side of the page is useful if the Base Pool contains endogenous interferents such as protein and triglycerides. The worksheet will calculate the amount of additional Interferent to add to reach the target concentration for testing.
2. The second page includes Replicate Tables to help determining the required number of replicates of Test and Control Pool for use with the Screening experiments. However, the Interference Screening Study worksheet will also calculate the number of replicates required, if Within-Run Standard Deviation and maximum allowable difference (Dmax) between Test and Control Pools are defined.
If in doubt, we suggest using a two-tailed test with an alpha of 5% (0.05) and a beta of 0.1 (10%).
3. For the Screening Experiment worksheets, enter the test results for Control and Test Pools, and the spreadsheet will calculate the difference, percent difference, upper and lower 95% confidence intervals for the difference, and a p-value for a two-sample t-test with equal variances. The t-test and 95% confidence intervals are provided for information purposes only; CLSI recommends interpreting the difference in terms of clinical significance—defined as a difference less than the Dmax.
Also, the spreadsheet will interpret the results for statistical and clinical significance.
The Worksheet includes a signature line and can be printed out to provide a summary of the experiments and the conclusions.
4. The Workbook also includes worksheets for dose-response experiments. Enter the Interferent concentrations of levels 1 and 5 and the intermediate concentrations will be calculated. Enter the test results and the means, difference, and percent difference will be calculated, in comparison to Interferent level 1.
The worksheet will also plot analyte concentration (Y) versus interference concentration (X) and calculate slope, intercept, and whether the slope is significantly different from zero. A slope of zero indicates that there is no association between analyte concentration and interferent concentration. This statistical test is provided for informational purposes only as CLSI recommends an interpretation based on clinical difference (Dmax) only.
Also, the worksheet will interpret concentration differences and/or percent differences depending on whether Dmax is defined as an absolute concentration or a percent difference. If Dmax is defined as both an absolute difference and percent difference the interpretation of clinical difference will require both to be greater than Dmax.
Finally, the worksheet will interpolate between each level of interferent to calculate the exact interferent concentration that is > Dmax. If multiple concentrations are provided, the correct concentration is the one that corresponds to the lowest interferent level that exceeds Dmax.
The Dose-Response worksheet also includes a signature line and can be printed out as a summary report.
1CLSI, Interference Testing in Clinical Chemistry; Approved Guidelines, 2ND ed, CLSI Document EP07-A2, Wayne, PA, 2005.
2Sonntag O, Scholer A. Drug interference in clinical chemistry: recommendations of drugs and their concentration to be used in drug interference studies. Ann Clin Biochem 2001; 38:376-385.

## Test Pool Calculator
TEST POOL CALCULATOR
| Date | Jan 25. 2019 |  | Units | mg/dL
| Interferent | Triglycerides |  | Tech | JHC
Use this calculator when the base pool concentration of interferent is negligible. | [5 empty] | Use this calculator when the base pool contains endogenous interferent. Recommended for Triglycerides, Protein, and most endogenous interferents.
Step 1: Determine Required Concentration of Interferent for Test Pool | [5 empty] | Step 1: Determine Concentration of Interferent in sample and in test pool
| Measured Interferent Concentration | Units | Dilution Factor Used for Measurement (B8) | Interferent Concentration | Units | Diluted Interferent Concentration | Units | Dilution Factor Used for Measurement (H7) | Interferent Concentration | Units |  | Endogenous Interferent Concentration (Base Pool) | Units
| 1000 | mg/dL | 15 | 15000 |  | 1000 | mg/dL | 15 | 15000 | mg/dL |  | 1000 | mg/dL
Step 2: Calculate Volumes for Test Pool | [5 empty] | Step 2: Calculate concentrated interferent and base pool addition
| Desired Test Pool Volume (mL) | Target Interferent Concentration | Units | Volume of Concentrated Interferent Required (mL) |  | Desired Test Pool Volume (mL) | Target Interferent Concentration | Units | Volume of Concentrated Interferent Required (mL)
| 10 | 1500 | mg/dL | 1 |  | 10 | 1500 | mg/dL | 0.33333
Step 3: Prepare Test Pool | [5 empty] | Step 3: Prepare Test Pool
Measure Base Pool into appropriate container, Total Volume: |  |  |  | 10 | mL | Measure Base Pool into appropriate container, Total Volume: |  |  |  | 10 | mL
Remove from base pool amount equal to volume of interferent to add |  |  |  | 1 | mL | Remove from base pool amount equal to volume of interferent to add |  |  |  | 0.33333 | mL
Add Interferent Stock |  |  |  | 1 | mL | Add Interferent Stock |  |  |  | 0.33333 | mL

## Replicate Tables
A reference for determining the number of replicates required for your Interference experiment.
One-sided Test with 5% Probability of Type 1 Error |  | Two-sided Test with 5% Probability of Type 1 Error |  | One-sided Test with 5% Probability of Type 1 Error |  | Two-sided Test with 5% Probability of Type 1 Error
and 10% Probability of Type II Error (90% Power) |  | and 10% Probability of Type II Error (90% Power) |  | and 5% Probability of Type II Error (95% Power) |  | and 5% Probability of Type II Error (95% Power)
dmax / s | No. of Replicates | dmax / s | No. of Replicates | dmax / s | No. of Replicates | dmax / s | No. of Replicates
0.8 | 27 | 0.8 | 33 | 0.8 | 34 | 0.8 | 41
1 | 18 | 1 | 22 | 1 | 22 | 1 | 26
1.1 | 15 | 1.1 | 18 | 1.1 | 18 | 1.1 | 22
1.2 | 12 | 1.2 | 15 | 1.2 | 16 | 1.2 | 19
1.3 | 11 | 1.3 | 13 | 1.3 | 13 | 1.3 | 16
1.4 | 9 | 1.4 | 11 | 1.4 | 12 | 1.4 | 14
1.5 | 8 | 1.5 | 10 | 1.5 | 10 | 1.5 | 12
1.6 | 7 | 1.6 | 9 | 1.6 | 9 | 1.6 | 11
1.8 | 6 | 1.8 | 7 | 1.8 | 7 | 1.8 | 9
2 | 5 | 2 | 6 | 2 | 6 | 2 | 7
2.5 | 3* | 2.5 | 4* | 2.5 | 4* | 2.5 | 5
3 | 2* | 3 | 3* | 3 | 3* | 3 | 3*
s = repeatability standard deviation (within-run SD) of the measurement procedure
dmax = maximum allowable difference to be detected at the analyte test concentration |  |  |  | One-sided Test with 5% Probability of Type 1 Error |  | Two-sided Test with 5% Probability of Type 1 Error
*A minimum of 5 repicates is recommended |  |  |  | and 20% Probability of Type II Error (80% Power) |  | and 20% Probability of Type II Error (80% Power)
|  |  |  | dmax / s | No. of Replicates | dmax / s | No. of Replicates
Note: Screening spreadsheet will also calculate sample size. |  |  |  | 0.8 | 20 | 0.8 | 25
|  |  |  | 1 | 13 | 1 | 16
|  |  |  | 1.1 | 11 | 1.1 | 13
|  |  |  | 1.2 | 9 | 1.2 | 11
|  |  |  | 1.3 | 8 | 1.3 | 10
|  |  |  | 1.4 | 7 | 1.4 | 9
|  |  |  | 1.5 | 6 | 1.5 | 7
|  |  |  | 1.6 | 5 | 1.6 | 7
|  |  |  | 1.8 | 4* | 1.8 | 5
|  |  |  | 2 | 4* | 2 | 4*
|  |  |  | 2.5 | 2* | 2.5 | 3*
|  |  |  | 3 | 2* | 3 | 2*

## Example-Screening
INTERFERENCE SCREENING STUDY |  |  |  |  | Enter Information and data in required fields (Yellow)
|  | *Enter required replicates or enter Dmax (cencentration) and Within-Run SD for estimate. | [7 empty] | 1-Sided | 2-Sided
Date: |  | 2012-05-11 |  |  | Test Substance |  |  | Hemolysate |  |  | α | 1-α | z1-α | 1-α | z1-α |  | β | 1-β | z1-β | Tails
Analyte |  | Creatinine |  |  | Estimated within-run SD |  |  | 0.1 |  | mg/dL | 0.2 | 0.8 | 0.842 | 0.8 | 1.282 |  | 0.2 | 0.8 | 0.842 | 1
Units |  | mg/dL |  |  | Dmax, % |  |  | 10% |  | % | 0.15 | 0.85 | 1.036 | 0.85 | 1.44 |  | 0.15 | 0.85 | 1.036 | 2
Method |  | Roche Jaffe |  |  | Dmax, concentration |  |  | 0.2 |  | mg/dL | 0.1 | 0.9 | 1.282 | 0.9 | 1.645 |  | 0.1 | 0.9 | 1.282
Reagent Lot# |  | 12345 |  |  | Alpha |  |  | 0.05 |  |  | 0.05 | 0.95 | 1.645 | 0.95 | 1.96 |  | 0.05 | 0.95 | 1.645
Calibrator Lot# |  | 98765 |  |  | Beta |  |  | 0.1 |  |  | 0.025 | 0.975 | 1.96 | 0.975 | 2.24 |  | 0.025 | 0.975 | 1.96
Instrument |  | Hitachi 911 |  |  | Replicates required, 1-tailed test |  |  | 5 |  |  | 0.01 | 0.99 | 2.326 | 0.99 | 2.576 |  | 0.01 | 0.99 | 2.326
[5 empty] | Replicates required, 2-tailed test |  |  | 6
Note: A minimum of 5 replicates is recommended. |  |  |  |  | Technician |  |  | JHC |  |  | Tails | Alpha | Beta | Number | Tails | Alpha | Beta | Number
Low Pool | [5 empty] | High Pool |  |  |  |  | 1 | 0.2 | 0.2 | 1.4179 | 2 | 0.2 | 0.2 | 2.2557
Replicate | Control | Replicate | Test |  |  | Replicate | Control | Replicate | Test |  | 1 | 0.2 | 0.15 | 1.7634 | 2 | 0.2 | 0.15 | 2.6866
C1 | 1.22 | T1 | 1.62 |  |  | C1 | 2.54 | T1 | 2.33 |  | 1 | 0.2 | 0.1 | 2.2557 | 2 | 0.2 | 0.1 | 3.287
C2 | 1.32 | T2 | 1.59 |  |  | C2 | 2.48 | T2 | 2.31 |  | 1 | 0.2 | 0.05 | 3.0926 | 2 | 0.2 | 0.05 | 4.2837
C3 | 1.18 | T3 | 1.58 |  |  | C3 | 2.46 | T3 | 2.41 |  | 1 | 0.2 | 0.025 | 3.9256 | 2 | 0.2 | 0.025 | 5.2553
C4 | 1.2 | T4 | 1.6 |  |  | C4 | 2.54 | T4 | 2.41 |  | 1 | 0.2 | 0.01 | 5.0181 | 2 | 0.2 | 0.01 | 6.5088
C5 | 1.21 | T5 | 1.59 |  |  | C5 | 2.52 | T5 | 2.33 |  | 1 | 0.15 | 0.2 | 1.7634 | 2 | 0.15 | 0.2 | 2.6038
C6 |  | T6 |  |  |  | C6 |  | T6 |  |  | 1 | 0.15 | 0.15 | 2.1466 | 2 | 0.15 | 0.15 | 3.0653
C7 |  | T7 |  |  |  | C7 |  | T7 |  |  | 1 | 0.15 | 0.1 | 2.6866 | 2 | 0.15 | 0.1 | 3.7046
C8 |  | T8 |  |  |  | C8 |  | T8 |  |  | 1 | 0.15 | 0.05 | 3.5939 | 2 | 0.15 | 0.05 | 4.7586
C9 |  | T9 |  |  |  | C9 |  | T9 |  |  | 1 | 0.15 | 0.025 | 4.488 | 2 | 0.15 | 0.025 | 5.78
C10 |  | T10 |  |  |  | C10 |  | T10 |  |  | 1 | 0.15 | 0.01 | 5.6515 | 2 | 0.15 | 0.01 | 7.0914
C11 |  | T11 |  |  |  | C11 |  | T11 |  |  | 1 | 0.1 | 0.2 | 2.2557 | 2 | 0.1 | 0.2 | 3.0926
C12 |  | T12 |  |  |  | C12 |  | T12 |  |  | 1 | 0.1 | 0.15 | 2.6866 | 2 | 0.1 | 0.15 | 3.5939
C13 |  | T13 |  |  |  | C13 |  | T13 |  |  | 1 | 0.1 | 0.1 | 3.287 | 2 | 0.1 | 0.1 | 4.2837
C14 |  | T14 |  |  |  | C14 |  | T14 |  |  | 1 | 0.1 | 0.05 | 4.2837 | 2 | 0.1 | 0.05 | 5.412
C15 |  | T15 |  |  |  | C15 |  | T15 |  |  | 1 | 0.1 | 0.025 | 5.2553 | 2 | 0.1 | 0.025 | 6.498
C16 |  | T16 |  |  |  | C16 |  | T16 |  |  | 1 | 0.1 | 0.01 | 6.5088 | 2 | 0.1 | 0.01 | 7.8844
C17 |  | T17 |  |  |  | C17 |  | T17 |  |  | 1 | 0.05 | 0.2 | 3.0926 | 2 | 0.05 | 0.2 | 3.9256
C18 |  | T18 |  |  |  | C18 |  | T18 |  |  | 1 | 0.05 | 0.15 | 3.5939 | 2 | 0.05 | 0.15 | 4.488
C19 |  | T19 |  |  |  | C19 |  | T19 |  |  | 1 | 0.05 | 0.1 | 4.2837 | 2 | 0.05 | 0.1 | 5.2553
C20 |  | T20 |  |  |  | C20 |  | T20 |  |  | 1 | 0.05 | 0.05 | 5.412 | 2 | 0.05 | 0.05 | 6.498
Count | 5 |  | 5 |  |  | Count | 5 |  | 5 |  | 1 | 0.05 | 0.025 | 6.498 | 2 | 0.05 | 0.025 | 7.6832
Mean | 1.226 |  | 1.596 |  |  | Mean | 2.508 |  | 2.358 |  | 1 | 0.05 | 0.01 | 7.8844 | 2 | 0.05 | 0.01 | 9.1849
SD | 0.054589 |  | 0.015166 |  |  | SD | 0.036332 |  | 0.048166 |  | 1 | 0.025 | 0.2 | 3.9256 | 2 | 0.025 | 0.2 | 4.7494
%CV | 4.453% |  | 0.9502% |  |  | %CV | 1.449% |  | 2.043% |  | 1 | 0.025 | 0.15 | 4.488 | 2 | 0.025 | 0.15 | 5.3661
p-value, 1-tailed, equal variance |  |  | 2.372e-07 |  |  | p-value, 1-tailed, equal variance |  |  | 0.00026753 |  | 1 | 0.025 | 0.1 | 5.2553 | 2 | 0.025 | 0.1 | 6.2022
p-value, 2-tailed, equal variance |  |  | 4.744e-07 |  |  | p-value, 2-tailed, equal variance |  |  | 0.00053506 |  | 1 | 0.025 | 0.05 | 6.498 | 2 | 0.025 | 0.05 | 7.5466
[11 empty] | 1 | 0.025 | 0.025 | 7.6832 | 2 | 0.025 | 0.025 | 8.82
Difference (Test - Control Means) |  |  |  | 0.37 |  | Difference (Test - Control Means) |  |  |  | -0.15 | 1 | 0.025 | 0.01 | 9.1849 | 2 | 0.025 | 0.01 | 10.4242
Percent Difference |  |  |  | 30.18% |  | Percent Difference |  |  |  | -5.981% | 1 | 0.01 | 0.2 | 5.0181 | 2 | 0.01 | 0.2 | 5.8414
Upper 95% Confidence Interval for the Difference |  |  |  | 0.42843 |  | Upper 95% Confidence Interval for the Difference |  |  |  | -0.087781 | 1 | 0.01 | 0.15 | 5.6515 | 2 | 0.01 | 0.15 | 6.5233
Lower 95% Confidence Interval for the Difference |  |  |  | 0.31157 |  | Lower 95% Confidence Interval for the Difference |  |  |  | -0.21222 | 1 | 0.01 | 0.1 | 6.5088 | 2 | 0.01 | 0.1 | 7.4421
[11 empty] | 1 | 0.01 | 0.05 | 7.8844 | 2 | 0.01 | 0.05 | 8.9084
Conclusions: Low Pool | [10 empty] | 1 | 0.01 | 0.025 | 9.1849 | 2 | 0.01 | 0.025 | 10.2876
The difference between Test and Control Pools IS statistically significant with a 1-tailed test. | [10 empty] | 1 | 0.01 | 0.01 | 10.8206 | 2 | 0.01 | 0.01 | 12.0148
The difference between Test and Control Pools IS statistically significant with a 2-tailed test.
The difference is > the acceptable limit (Dmax), indicating that Interference is Clinically Significant
Low Pool
Conclusions: High Pool | [11 empty] | alpha | 0.05
The difference between Test and Control Pools IS statistically significant with a 1-tailed test. | [11 empty] | DF | 8 |  | *=B34+D34-2
The difference between Test and Control Pools IS statistically significant with a 2-tailed test. | [11 empty] | t-value | 2.306 |  | *=T.INV.2T(O51,52)
The difference is < the acceptable limit (Dmax), indicating that Interference is NOT Clinically Significant | [12 empty] | 0.025338 |  | *=SQRT(((B36^2)/B34)+((D36^2)/D34))
[12 empty] | 95% CI | 0.058429 |  | *=O53*O54
APPROVED: |  |  |  |  | Date: | [6 empty] | High Pool
[12 empty] | alpha | 0.05
[12 empty] | DF | 8 |  | *=H34+J34-2
Name: | [11 empty] | t-value | 2.306 |  | *=T.INV.2T(O58,O59)
[13 empty] | 0.026981 |  | *=SQRT(((H36^2)/H34)+((J36^2)/J34))
Notes: | [11 empty] | 95% CI | 0.062219 |  | *=O60*O61
1 | If there are missing data points (less than required replicates) repeat the experiment.
2 | Review data for outliers and exclude data with an assignable cause. If there is more than one outlier consider repeating the experiment.
3 | Conclusion based on statistical significance is provided for informational purposes. CLSI EP7 recommends interpretation based on clinical utility.
INTERFERENCE SCREENING
EXAMPLE

## Example-Dose Response
| FIVE LEVEL INTERFERENCE DOSE RESPONSE STUDY |  |  |  | *Enter Dmax as absolute concentration, percent difference or both
ANALYTE 1
| Date: | Sept. 6, 2018 |  |  | Test Substance (Interferent) |  |  | Triglycerides (Lipemia)
| Analyte | Glucose |  |  | Technician |  |  | JHC
| Units | mg/dL |  |  | *Maximum Allowable Difference (Dmax), Concentration |  |  | 5 |  | mg/dL
| Method | Roche HK |  |  | *Maximum Allowable Difference (Dmax), Percent |  |  | 10% |  | %
| Reagent Lot# | 12345
| Calibrator Lot# | 987654
| Instrument | Cobas
Enter concentrations of Levels 1 and 5, and intermIxed levels will be calculated; Otherwise, enter concentrations for all five levels.
Interferent | Interferent | Rep 1 | Rep 2 | Rep 3 | Rep 4 | Rep 5 | Mean | Absolute | Absolute | Percent | Percent
Level | Concentration | Concentration | Concentration | Concentration | Concentration | Concentration | Concentration | Difference | Limit | Difference | Limit
1 | 0 | 120 | 116 | 118 | 121 | 116 | 118.2 | 0 | Exceeded? | 0% | Exceeded?
2 | 250 | 122 | 122 | 121 | 120 | 123 | 121.6 | 3.4 | NO | 2.876% | NO
3 | 500 | 118 | 123 | 121 | 125 | 118 | 121 | 2.8 | NO | 2.369% | NO
4 | 750 | 125 | 126 | 128 | 122 | 118 | 123.8 | 5.6 | YES | 4.738% | NO
5 | 1000 | 130 | 134 | 136 | 131 | 133 | 132.8 | 14.6 | YES | 12.35% | YES
Mean X | 500
Mean Y | 123.48 | [7 empty] | Xi-Xm^2 |  | Yi-Ym^2
Intercept | 117.2 | [6 empty] | X1 | 250000 | Y1 | 27.8784
Slope | 0.01256 | [6 empty] | X2 | 62500 | Y2 | 3.5344
r | 0.88981 | [6 empty] | X3 | 0 | Y3 | 6.1504
r2 | 0.79176 | [6 empty] | X4 | 62500 | Y4 | 0.1024
Sy/x | 2.9401 | [6 empty] | X5 | 250000 | Y5 | 86.8624
Sum (Xi-Xm)^2 | 625000
Sum (Yi-Ym)^2 | 124.528 | [7 empty] | Control Line | (+) Dmax | (-) Dmax
SE(slope) | 0.0029433 | [6 empty] | X1 | 0 | 0 | 0
N | 25 | [6 empty] | X5 | 1000 | 1000 | 1000
t-statistic | 4.2674 | [6 empty] | Y1 | 118.2 | 123.2 | 113.2
DF | 23 | [6 empty] | Y5 | 118.2 | 123.2 | 113.2
t-Comparison | 2.0687
Slope p≤0.05? | YES | [7 empty] | Control Line | (+)%Dmax | (-)%Dmax
[8 empty] | X1 | 0 | 0 | 0
Conclusions: Dose-Response Experiment | [7 empty] | X5 | 1000 | 1000 | 1000
The slope is NOT statistically significantly different from zero. | [7 empty] | Y1 | 118.2 | 130.02 | 106.38
[8 empty] | Y5 | 118.2 | 130.02 | 106.38
Bias with Interferent Level 2 < the acceptable limit (Dmax), indicating that Interference is NOT Clinically Significant.
Bias with Interferent Level 3 < the acceptable limit (Dmax), indicating that Interference is NOT Clinically Significant.
Bias with Interferent Level 4 < the acceptable limit (Dmax), indicating that Interference is NOT Clinically Significant.
Bias with Interferent Level 5 > the acceptable limit (Dmax), indicating that Interference IS Clinically Significant.
POINT-TO-POINT INTERFERENCE CUTPOINT ESTIMATE | [8 empty] | Greater
Absolute Difference | Slope | Intercept | Interference |  | Percent Difference | Slope | Intercept | Interference | Abs or %
Level 1 to Level 2 | 0.0136 | 0 | 367.6471 |  | Level 1 to Level 2 | 0.00011506 | 1.7347e-18 | 869.1176 | 869.1176
Level 2 to Level 3 | -0.0024 | 4 | 416.6667 |  | Level 2 to Level 3 | -2.0305e-05 | 0.033841 | 3258.3333 | 3258.3333
Level 3 to Level 4 | 0.0112 | -2.8 | 696.4286 |  | Level 3 to Level 4 | 9.4755e-05 | -0.023689 | 1305.3571 | 1305.3571
Level 4 to Level 5 | 0.036 | -21.4 | 733.3333 |  | Level 4 to Level 5 | 0.00030457 | -0.18105 | 922.7778 | 922.7778
INTERPOLATED: |  | Level 1-2 | Level 2-3 | Level 3-4 | Level 4-5
Allowable Interferent Concentration |  |  |  |  | 922.7778
APPROVED:
Name: |  |  |  |  | Date:
Notes:
1 | If there are more than two (2) missing data points (out of 25) consider repeating the experiment.
Alternatively, generate additional measurements to replace missing values; understanding that total measurement variability may increase.
2 | Review data for outliers and exclude data with an assignable cause. If there is more than one outlier consider repeating the experiment.
3 | Conclusion based on statistical significance is provided for informational purposes. CLSI EP7 recommends interpretation based on clinical utility.
INTERFERENCE DOSE RESPONSE
EXAMPLE

## Screening Worksheet 1
INTERFERENCE SCREENING STUDY |  |  |  |  | Enter Information and data in required fields (Yellow)
|  | *Enter required replicates or enter Dmax (cencentration) and Within-Run SD for estimate. | [7 empty] | 1-Sided | 2-Sided
Date: |  |  |  |  | Test Substance | [5 empty] | α | 1-α | z1-α | 1-α | z1-α |  | β | 1-β | z1-β | Tails
Analyte |  |  |  |  | Estimated within-run SD |  |  |  |  | 0 | 0.2 | 0.8 | 0.842 | 0.8 | 1.282 |  | 0.2 | 0.8 | 0.842 | 1
Units |  |  |  |  | Dmax, % |  |  |  |  | % | 0.15 | 0.85 | 1.036 | 0.85 | 1.44 |  | 0.15 | 0.85 | 1.036 | 2
Method |  |  |  |  | Dmax, concentration |  |  |  |  | 0 | 0.1 | 0.9 | 1.282 | 0.9 | 1.645 |  | 0.1 | 0.9 | 1.282
Reagent Lot# |  |  |  |  | Alpha | [5 empty] | 0.05 | 0.95 | 1.645 | 0.95 | 1.96 |  | 0.05 | 0.95 | 1.645
Calibrator Lot# |  |  |  |  | Beta | [5 empty] | 0.025 | 0.975 | 1.96 | 0.975 | 2.24 |  | 0.025 | 0.975 | 1.96
Instrument |  |  |  |  | Replicates required, 1-tailed test |  |  | 0 |  |  | 0.01 | 0.99 | 2.326 | 0.99 | 2.576 |  | 0.01 | 0.99 | 2.326
[5 empty] | Replicates required, 2-tailed test |  |  | 0
Note: A minimum of 5 replicates is recommended. |  |  |  |  | Technician | [5 empty] | Tails | Alpha | Beta | Number | Tails | Alpha | Beta | Number
Low Pool | [5 empty] | High Pool |  |  |  |  | 1 | 0.2 | 0.2 | #DIV/0! | 2 | 0.2 | 0.2 | #DIV/0!
Replicate | Control | Replicate | Test |  |  | Replicate | Control | Replicate | Test |  | 1 | 0.2 | 0.15 | #DIV/0! | 2 | 0.2 | 0.15 | #DIV/0!
C1 |  | T1 |  |  |  | C1 |  | T1 |  |  | 1 | 0.2 | 0.1 | #DIV/0! | 2 | 0.2 | 0.1 | #DIV/0!
C2 |  | T2 |  |  |  | C2 |  | T2 |  |  | 1 | 0.2 | 0.05 | #DIV/0! | 2 | 0.2 | 0.05 | #DIV/0!
C3 |  | T3 |  |  |  | C3 |  | T3 |  |  | 1 | 0.2 | 0.025 | #DIV/0! | 2 | 0.2 | 0.025 | #DIV/0!
C4 |  | T4 |  |  |  | C4 |  | T4 |  |  | 1 | 0.2 | 0.01 | #DIV/0! | 2 | 0.2 | 0.01 | #DIV/0!
C5 |  | T5 |  |  |  | C5 |  | T5 |  |  | 1 | 0.15 | 0.2 | #DIV/0! | 2 | 0.15 | 0.2 | #DIV/0!
C6 |  | T6 |  |  |  | C6 |  | T6 |  |  | 1 | 0.15 | 0.15 | #DIV/0! | 2 | 0.15 | 0.15 | #DIV/0!
C7 |  | T7 |  |  |  | C7 |  | T7 |  |  | 1 | 0.15 | 0.1 | #DIV/0! | 2 | 0.15 | 0.1 | #DIV/0!
C8 |  | T8 |  |  |  | C8 |  | T8 |  |  | 1 | 0.15 | 0.05 | #DIV/0! | 2 | 0.15 | 0.05 | #DIV/0!
C9 |  | T9 |  |  |  | C9 |  | T9 |  |  | 1 | 0.15 | 0.025 | #DIV/0! | 2 | 0.15 | 0.025 | #DIV/0!
C10 |  | T10 |  |  |  | C10 |  | T10 |  |  | 1 | 0.15 | 0.01 | #DIV/0! | 2 | 0.15 | 0.01 | #DIV/0!
C11 |  | T11 |  |  |  | C11 |  | T11 |  |  | 1 | 0.1 | 0.2 | #DIV/0! | 2 | 0.1 | 0.2 | #DIV/0!
C12 |  | T12 |  |  |  | C12 |  | T12 |  |  | 1 | 0.1 | 0.15 | #DIV/0! | 2 | 0.1 | 0.15 | #DIV/0!
C13 |  | T13 |  |  |  | C13 |  | T13 |  |  | 1 | 0.1 | 0.1 | #DIV/0! | 2 | 0.1 | 0.1 | #DIV/0!
C14 |  | T14 |  |  |  | C14 |  | T14 |  |  | 1 | 0.1 | 0.05 | #DIV/0! | 2 | 0.1 | 0.05 | #DIV/0!
C15 |  | T15 |  |  |  | C15 |  | T15 |  |  | 1 | 0.1 | 0.025 | #DIV/0! | 2 | 0.1 | 0.025 | #DIV/0!
C16 |  | T16 |  |  |  | C16 |  | T16 |  |  | 1 | 0.1 | 0.01 | #DIV/0! | 2 | 0.1 | 0.01 | #DIV/0!
C17 |  | T17 |  |  |  | C17 |  | T17 |  |  | 1 | 0.05 | 0.2 | #DIV/0! | 2 | 0.05 | 0.2 | #DIV/0!
C18 |  | T18 |  |  |  | C18 |  | T18 |  |  | 1 | 0.05 | 0.15 | #DIV/0! | 2 | 0.05 | 0.15 | #DIV/0!
C19 |  | T19 |  |  |  | C19 |  | T19 |  |  | 1 | 0.05 | 0.1 | #DIV/0! | 2 | 0.05 | 0.1 | #DIV/0!
C20 |  | T20 |  |  |  | C20 |  | T20 |  |  | 1 | 0.05 | 0.05 | #DIV/0! | 2 | 0.05 | 0.05 | #DIV/0!
Count | 0 |  | 0 |  |  | Count | 0 |  | 0 |  | 1 | 0.05 | 0.025 | #DIV/0! | 2 | 0.05 | 0.025 | #DIV/0!
Mean | #DIV/0! |  | #DIV/0! |  |  | Mean | #DIV/0! |  | #DIV/0! |  | 1 | 0.05 | 0.01 | #DIV/0! | 2 | 0.05 | 0.01 | #DIV/0!
SD | #DIV/0! |  | #DIV/0! |  |  | SD | #DIV/0! |  | #DIV/0! |  | 1 | 0.025 | 0.2 | #DIV/0! | 2 | 0.025 | 0.2 | #DIV/0!
%CV | #DIV/0! |  | #DIV/0! |  |  | %CV | #DIV/0! |  | #DIV/0! |  | 1 | 0.025 | 0.15 | #DIV/0! | 2 | 0.025 | 0.15 | #DIV/0!
p-value, 1-tailed, equal variance |  |  | #DIV/0! |  |  | p-value, 1-tailed, equal variance |  |  | #DIV/0! |  | 1 | 0.025 | 0.1 | #DIV/0! | 2 | 0.025 | 0.1 | #DIV/0!
p-value, 2-tailed, equal variance |  |  | #DIV/0! |  |  | p-value, 2-tailed, equal variance |  |  | #DIV/0! |  | 1 | 0.025 | 0.05 | #DIV/0! | 2 | 0.025 | 0.05 | #DIV/0!
[11 empty] | 1 | 0.025 | 0.025 | #DIV/0! | 2 | 0.025 | 0.025 | #DIV/0!
Difference (Test - Control Means) |  |  |  | #DIV/0! |  | Difference (Test - Control Means) |  |  |  | #DIV/0! | 1 | 0.025 | 0.01 | #DIV/0! | 2 | 0.025 | 0.01 | #DIV/0!
Percent Difference |  |  |  | #DIV/0! |  | Percent Difference |  |  |  | #DIV/0! | 1 | 0.01 | 0.2 | #DIV/0! | 2 | 0.01 | 0.2 | #DIV/0!
Upper 95% Confidence Interval for the Difference |  |  |  | #DIV/0! |  | Upper 95% Confidence Interval for the Difference |  |  |  | #DIV/0! | 1 | 0.01 | 0.15 | #DIV/0! | 2 | 0.01 | 0.15 | #DIV/0!
Lower 95% Confidence Interval for the Difference |  |  |  | #DIV/0! |  | Lower 95% Confidence Interval for the Difference |  |  |  | #DIV/0! | 1 | 0.01 | 0.1 | #DIV/0! | 2 | 0.01 | 0.1 | #DIV/0!
[11 empty] | 1 | 0.01 | 0.05 | #DIV/0! | 2 | 0.01 | 0.05 | #DIV/0!
Conclusions: Low Pool | [10 empty] | 1 | 0.01 | 0.025 | #DIV/0! | 2 | 0.01 | 0.025 | #DIV/0!
#DIV/0! | [10 empty] | 1 | 0.01 | 0.01 | #DIV/0! | 2 | 0.01 | 0.01 | #DIV/0!
#DIV/0!
#DIV/0!
Low Pool
Conclusions: High Pool | [11 empty] | alpha | 0.05
#DIV/0! | [11 empty] | DF | -2 |  | *=B34+D34-2
#DIV/0! | [11 empty] | t-value | #NUM! |  | *=T.INV.2T(O51,52)
#DIV/0! | [12 empty] | #DIV/0! |  | *=SQRT(((B36^2)/B34)+((D36^2)/D34))
[12 empty] | 95% CI | #NUM! |  | *=O53*O54
APPROVED: |  |  |  |  | Date: | [6 empty] | High Pool
[12 empty] | alpha | 0.05
[12 empty] | DF | -2 |  | *=H34+J34-2
Name: | [11 empty] | t-value | #NUM! |  | *=T.INV.2T(O58,O59)
[13 empty] | #DIV/0! |  | *=SQRT(((H36^2)/H34)+((J36^2)/J34))
Notes: | [11 empty] | 95% CI | #NUM! |  | *=O60*O61
1 | If there are missing data points (less than required replicates) repeat the experiment.
2 | Review data for outliers and exclude data with an assignable cause. If there is more than one outlier consider repeating the experiment.
3 | Conclusion based on statistical significance is provided for informational purposes. CLSI EP7 recommends interpretation based on clinical utility.
INTERFERENCE SCREENING
EXAMPLE

## Screening Worksheet 2
INTERFERENCE SCREENING STUDY |  |  |  |  | Enter Information and data in required fields (Yellow)
|  | *Enter required replicates or enter Dmax (cencentration) and Within-Run SD for estimate. | [7 empty] | 1-Sided | 2-Sided
Date: |  |  |  |  | Test Substance | [5 empty] | α | 1-α | z1-α | 1-α | z1-α |  | β | 1-β | z1-β | Tails
Analyte |  |  |  |  | Estimated within-run SD |  |  |  |  | 0 | 0.2 | 0.8 | 0.842 | 0.8 | 1.282 |  | 0.2 | 0.8 | 0.842 | 1
Units |  |  |  |  | Dmax, % |  |  |  |  | % | 0.15 | 0.85 | 1.036 | 0.85 | 1.44 |  | 0.15 | 0.85 | 1.036 | 2
Method |  |  |  |  | Dmax, concentration |  |  |  |  | 0 | 0.1 | 0.9 | 1.282 | 0.9 | 1.645 |  | 0.1 | 0.9 | 1.282
Reagent Lot# |  |  |  |  | Alpha | [5 empty] | 0.05 | 0.95 | 1.645 | 0.95 | 1.96 |  | 0.05 | 0.95 | 1.645
Calibrator Lot# |  |  |  |  | Beta | [5 empty] | 0.025 | 0.975 | 1.96 | 0.975 | 2.24 |  | 0.025 | 0.975 | 1.96
Instrument |  |  |  |  | Replicates required, 1-tailed test |  |  | 0 |  |  | 0.01 | 0.99 | 2.326 | 0.99 | 2.576 |  | 0.01 | 0.99 | 2.326
[5 empty] | Replicates required, 2-tailed test |  |  | 0
Note: A minimum of 5 replicates is recommended. |  |  |  |  | Technician | [5 empty] | Tails | Alpha | Beta | Number | Tails | Alpha | Beta | Number
Low Pool | [5 empty] | High Pool |  |  |  |  | 1 | 0.2 | 0.2 | #DIV/0! | 2 | 0.2 | 0.2 | #DIV/0!
Replicate | Control | Replicate | Test |  |  | Replicate | Control | Replicate | Test |  | 1 | 0.2 | 0.15 | #DIV/0! | 2 | 0.2 | 0.15 | #DIV/0!
C1 |  | T1 |  |  |  | C1 |  | T1 |  |  | 1 | 0.2 | 0.1 | #DIV/0! | 2 | 0.2 | 0.1 | #DIV/0!
C2 |  | T2 |  |  |  | C2 |  | T2 |  |  | 1 | 0.2 | 0.05 | #DIV/0! | 2 | 0.2 | 0.05 | #DIV/0!
C3 |  | T3 |  |  |  | C3 |  | T3 |  |  | 1 | 0.2 | 0.025 | #DIV/0! | 2 | 0.2 | 0.025 | #DIV/0!
C4 |  | T4 |  |  |  | C4 |  | T4 |  |  | 1 | 0.2 | 0.01 | #DIV/0! | 2 | 0.2 | 0.01 | #DIV/0!
C5 |  | T5 |  |  |  | C5 |  | T5 |  |  | 1 | 0.15 | 0.2 | #DIV/0! | 2 | 0.15 | 0.2 | #DIV/0!
C6 |  | T6 |  |  |  | C6 |  | T6 |  |  | 1 | 0.15 | 0.15 | #DIV/0! | 2 | 0.15 | 0.15 | #DIV/0!
C7 |  | T7 |  |  |  | C7 |  | T7 |  |  | 1 | 0.15 | 0.1 | #DIV/0! | 2 | 0.15 | 0.1 | #DIV/0!
C8 |  | T8 |  |  |  | C8 |  | T8 |  |  | 1 | 0.15 | 0.05 | #DIV/0! | 2 | 0.15 | 0.05 | #DIV/0!
C9 |  | T9 |  |  |  | C9 |  | T9 |  |  | 1 | 0.15 | 0.025 | #DIV/0! | 2 | 0.15 | 0.025 | #DIV/0!
C10 |  | T10 |  |  |  | C10 |  | T10 |  |  | 1 | 0.15 | 0.01 | #DIV/0! | 2 | 0.15 | 0.01 | #DIV/0!
C11 |  | T11 |  |  |  | C11 |  | T11 |  |  | 1 | 0.1 | 0.2 | #DIV/0! | 2 | 0.1 | 0.2 | #DIV/0!
C12 |  | T12 |  |  |  | C12 |  | T12 |  |  | 1 | 0.1 | 0.15 | #DIV/0! | 2 | 0.1 | 0.15 | #DIV/0!
C13 |  | T13 |  |  |  | C13 |  | T13 |  |  | 1 | 0.1 | 0.1 | #DIV/0! | 2 | 0.1 | 0.1 | #DIV/0!
C14 |  | T14 |  |  |  | C14 |  | T14 |  |  | 1 | 0.1 | 0.05 | #DIV/0! | 2 | 0.1 | 0.05 | #DIV/0!
C15 |  | T15 |  |  |  | C15 |  | T15 |  |  | 1 | 0.1 | 0.025 | #DIV/0! | 2 | 0.1 | 0.025 | #DIV/0!
C16 |  | T16 |  |  |  | C16 |  | T16 |  |  | 1 | 0.1 | 0.01 | #DIV/0! | 2 | 0.1 | 0.01 | #DIV/0!
C17 |  | T17 |  |  |  | C17 |  | T17 |  |  | 1 | 0.05 | 0.2 | #DIV/0! | 2 | 0.05 | 0.2 | #DIV/0!
C18 |  | T18 |  |  |  | C18 |  | T18 |  |  | 1 | 0.05 | 0.15 | #DIV/0! | 2 | 0.05 | 0.15 | #DIV/0!
C19 |  | T19 |  |  |  | C19 |  | T19 |  |  | 1 | 0.05 | 0.1 | #DIV/0! | 2 | 0.05 | 0.1 | #DIV/0!
C20 |  | T20 |  |  |  | C20 |  | T20 |  |  | 1 | 0.05 | 0.05 | #DIV/0! | 2 | 0.05 | 0.05 | #DIV/0!
Count | 0 |  | 0 |  |  | Count | 0 |  | 0 |  | 1 | 0.05 | 0.025 | #DIV/0! | 2 | 0.05 | 0.025 | #DIV/0!
Mean | #DIV/0! |  | #DIV/0! |  |  | Mean | #DIV/0! |  | #DIV/0! |  | 1 | 0.05 | 0.01 | #DIV/0! | 2 | 0.05 | 0.01 | #DIV/0!
SD | #DIV/0! |  | #DIV/0! |  |  | SD | #DIV/0! |  | #DIV/0! |  | 1 | 0.025 | 0.2 | #DIV/0! | 2 | 0.025 | 0.2 | #DIV/0!
%CV | #DIV/0! |  | #DIV/0! |  |  | %CV | #DIV/0! |  | #DIV/0! |  | 1 | 0.025 | 0.15 | #DIV/0! | 2 | 0.025 | 0.15 | #DIV/0!
p-value, 1-tailed, equal variance |  |  | #DIV/0! |  |  | p-value, 1-tailed, equal variance |  |  | #DIV/0! |  | 1 | 0.025 | 0.1 | #DIV/0! | 2 | 0.025 | 0.1 | #DIV/0!
p-value, 2-tailed, equal variance |  |  | #DIV/0! |  |  | p-value, 2-tailed, equal variance |  |  | #DIV/0! |  | 1 | 0.025 | 0.05 | #DIV/0! | 2 | 0.025 | 0.05 | #DIV/0!
[11 empty] | 1 | 0.025 | 0.025 | #DIV/0! | 2 | 0.025 | 0.025 | #DIV/0!
Difference (Test - Control Means) |  |  |  | #DIV/0! |  | Difference (Test - Control Means) |  |  |  | #DIV/0! | 1 | 0.025 | 0.01 | #DIV/0! | 2 | 0.025 | 0.01 | #DIV/0!
Percent Difference |  |  |  | #DIV/0! |  | Percent Difference |  |  |  | #DIV/0! | 1 | 0.01 | 0.2 | #DIV/0! | 2 | 0.01 | 0.2 | #DIV/0!
Upper 95% Confidence Interval for the Difference |  |  |  | #DIV/0! |  | Upper 95% Confidence Interval for the Difference |  |  |  | #DIV/0! | 1 | 0.01 | 0.15 | #DIV/0! | 2 | 0.01 | 0.15 | #DIV/0!
Lower 95% Confidence Interval for the Difference |  |  |  | #DIV/0! |  | Lower 95% Confidence Interval for the Difference |  |  |  | #DIV/0! | 1 | 0.01 | 0.1 | #DIV/0! | 2 | 0.01 | 0.1 | #DIV/0!
[11 empty] | 1 | 0.01 | 0.05 | #DIV/0! | 2 | 0.01 | 0.05 | #DIV/0!
Conclusions: Low Pool | [10 empty] | 1 | 0.01 | 0.025 | #DIV/0! | 2 | 0.01 | 0.025 | #DIV/0!
#DIV/0! | [10 empty] | 1 | 0.01 | 0.01 | #DIV/0! | 2 | 0.01 | 0.01 | #DIV/0!
#DIV/0!
#DIV/0!
Low Pool
Conclusions: High Pool | [11 empty] | alpha | 0.05
#DIV/0! | [11 empty] | DF | -2 |  | *=B34+D34-2
#DIV/0! | [11 empty] | t-value | #NUM! |  | *=T.INV.2T(O51,52)
#DIV/0! | [12 empty] | #DIV/0! |  | *=SQRT(((B36^2)/B34)+((D36^2)/D34))
[12 empty] | 95% CI | #NUM! |  | *=O53*O54
APPROVED: |  |  |  |  | Date: | [6 empty] | High Pool
[12 empty] | alpha | 0.05
[12 empty] | DF | -2 |  | *=H34+J34-2
Name: | [11 empty] | t-value | #NUM! |  | *=T.INV.2T(O58,O59)
[13 empty] | #DIV/0! |  | *=SQRT(((H36^2)/H34)+((J36^2)/J34))
Notes: | [11 empty] | 95% CI | #NUM! |  | *=O60*O61
1 | If there are missing data points (less than required replicates) repeat the experiment.
2 | Review data for outliers and exclude data with an assignable cause. If there is more than one outlier consider repeating the experiment.
3 | Conclusion based on statistical significance is provided for informational purposes. CLSI EP7 recommends interpretation based on clinical utility.
INTERFERENCE SCREENING
EXAMPLE

## Screening Worksheet 3
INTERFERENCE SCREENING STUDY |  |  |  |  | Enter Information and data in required fields (Yellow)
|  | *Enter required replicates or enter Dmax (cencentration) and Within-Run SD for estimate. | [7 empty] | 1-Sided | 2-Sided
Date: |  |  |  |  | Test Substance | [5 empty] | α | 1-α | z1-α | 1-α | z1-α |  | β | 1-β | z1-β | Tails
Analyte |  |  |  |  | Estimated within-run SD |  |  |  |  | 0 | 0.2 | 0.8 | 0.842 | 0.8 | 1.282 |  | 0.2 | 0.8 | 0.842 | 1
Units |  |  |  |  | Dmax, % |  |  |  |  | % | 0.15 | 0.85 | 1.036 | 0.85 | 1.44 |  | 0.15 | 0.85 | 1.036 | 2
Method |  |  |  |  | Dmax, concentration |  |  |  |  | 0 | 0.1 | 0.9 | 1.282 | 0.9 | 1.645 |  | 0.1 | 0.9 | 1.282
Reagent Lot# |  |  |  |  | Alpha | [5 empty] | 0.05 | 0.95 | 1.645 | 0.95 | 1.96 |  | 0.05 | 0.95 | 1.645
Calibrator Lot# |  |  |  |  | Beta | [5 empty] | 0.025 | 0.975 | 1.96 | 0.975 | 2.24 |  | 0.025 | 0.975 | 1.96
Instrument |  |  |  |  | Replicates required, 1-tailed test |  |  | 0 |  |  | 0.01 | 0.99 | 2.326 | 0.99 | 2.576 |  | 0.01 | 0.99 | 2.326
[5 empty] | Replicates required, 2-tailed test |  |  | 0
Note: A minimum of 5 replicates is recommended. |  |  |  |  | Technician | [5 empty] | Tails | Alpha | Beta | Number | Tails | Alpha | Beta | Number
Low Pool | [5 empty] | High Pool |  |  |  |  | 1 | 0.2 | 0.2 | #DIV/0! | 2 | 0.2 | 0.2 | #DIV/0!
Replicate | Control | Replicate | Test |  |  | Replicate | Control | Replicate | Test |  | 1 | 0.2 | 0.15 | #DIV/0! | 2 | 0.2 | 0.15 | #DIV/0!
C1 |  | T1 |  |  |  | C1 |  | T1 |  |  | 1 | 0.2 | 0.1 | #DIV/0! | 2 | 0.2 | 0.1 | #DIV/0!
C2 |  | T2 |  |  |  | C2 |  | T2 |  |  | 1 | 0.2 | 0.05 | #DIV/0! | 2 | 0.2 | 0.05 | #DIV/0!
C3 |  | T3 |  |  |  | C3 |  | T3 |  |  | 1 | 0.2 | 0.025 | #DIV/0! | 2 | 0.2 | 0.025 | #DIV/0!
C4 |  | T4 |  |  |  | C4 |  | T4 |  |  | 1 | 0.2 | 0.01 | #DIV/0! | 2 | 0.2 | 0.01 | #DIV/0!
C5 |  | T5 |  |  |  | C5 |  | T5 |  |  | 1 | 0.15 | 0.2 | #DIV/0! | 2 | 0.15 | 0.2 | #DIV/0!
C6 |  | T6 |  |  |  | C6 |  | T6 |  |  | 1 | 0.15 | 0.15 | #DIV/0! | 2 | 0.15 | 0.15 | #DIV/0!
C7 |  | T7 |  |  |  | C7 |  | T7 |  |  | 1 | 0.15 | 0.1 | #DIV/0! | 2 | 0.15 | 0.1 | #DIV/0!
C8 |  | T8 |  |  |  | C8 |  | T8 |  |  | 1 | 0.15 | 0.05 | #DIV/0! | 2 | 0.15 | 0.05 | #DIV/0!
C9 |  | T9 |  |  |  | C9 |  | T9 |  |  | 1 | 0.15 | 0.025 | #DIV/0! | 2 | 0.15 | 0.025 | #DIV/0!
C10 |  | T10 |  |  |  | C10 |  | T10 |  |  | 1 | 0.15 | 0.01 | #DIV/0! | 2 | 0.15 | 0.01 | #DIV/0!
C11 |  | T11 |  |  |  | C11 |  | T11 |  |  | 1 | 0.1 | 0.2 | #DIV/0! | 2 | 0.1 | 0.2 | #DIV/0!
C12 |  | T12 |  |  |  | C12 |  | T12 |  |  | 1 | 0.1 | 0.15 | #DIV/0! | 2 | 0.1 | 0.15 | #DIV/0!
C13 |  | T13 |  |  |  | C13 |  | T13 |  |  | 1 | 0.1 | 0.1 | #DIV/0! | 2 | 0.1 | 0.1 | #DIV/0!
C14 |  | T14 |  |  |  | C14 |  | T14 |  |  | 1 | 0.1 | 0.05 | #DIV/0! | 2 | 0.1 | 0.05 | #DIV/0!
C15 |  | T15 |  |  |  | C15 |  | T15 |  |  | 1 | 0.1 | 0.025 | #DIV/0! | 2 | 0.1 | 0.025 | #DIV/0!
C16 |  | T16 |  |  |  | C16 |  | T16 |  |  | 1 | 0.1 | 0.01 | #DIV/0! | 2 | 0.1 | 0.01 | #DIV/0!
C17 |  | T17 |  |  |  | C17 |  | T17 |  |  | 1 | 0.05 | 0.2 | #DIV/0! | 2 | 0.05 | 0.2 | #DIV/0!
C18 |  | T18 |  |  |  | C18 |  | T18 |  |  | 1 | 0.05 | 0.15 | #DIV/0! | 2 | 0.05 | 0.15 | #DIV/0!
C19 |  | T19 |  |  |  | C19 |  | T19 |  |  | 1 | 0.05 | 0.1 | #DIV/0! | 2 | 0.05 | 0.1 | #DIV/0!
C20 |  | T20 |  |  |  | C20 |  | T20 |  |  | 1 | 0.05 | 0.05 | #DIV/0! | 2 | 0.05 | 0.05 | #DIV/0!
Count | 0 |  | 0 |  |  | Count | 0 |  | 0 |  | 1 | 0.05 | 0.025 | #DIV/0! | 2 | 0.05 | 0.025 | #DIV/0!
Mean | #DIV/0! |  | #DIV/0! |  |  | Mean | #DIV/0! |  | #DIV/0! |  | 1 | 0.05 | 0.01 | #DIV/0! | 2 | 0.05 | 0.01 | #DIV/0!
SD | #DIV/0! |  | #DIV/0! |  |  | SD | #DIV/0! |  | #DIV/0! |  | 1 | 0.025 | 0.2 | #DIV/0! | 2 | 0.025 | 0.2 | #DIV/0!
%CV | #DIV/0! |  | #DIV/0! |  |  | %CV | #DIV/0! |  | #DIV/0! |  | 1 | 0.025 | 0.15 | #DIV/0! | 2 | 0.025 | 0.15 | #DIV/0!
p-value, 1-tailed, equal variance |  |  | #DIV/0! |  |  | p-value, 1-tailed, equal variance |  |  | #DIV/0! |  | 1 | 0.025 | 0.1 | #DIV/0! | 2 | 0.025 | 0.1 | #DIV/0!
p-value, 2-tailed, equal variance |  |  | #DIV/0! |  |  | p-value, 2-tailed, equal variance |  |  | #DIV/0! |  | 1 | 0.025 | 0.05 | #DIV/0! | 2 | 0.025 | 0.05 | #DIV/0!
[11 empty] | 1 | 0.025 | 0.025 | #DIV/0! | 2 | 0.025 | 0.025 | #DIV/0!
Difference (Test - Control Means) |  |  |  | #DIV/0! |  | Difference (Test - Control Means) |  |  |  | #DIV/0! | 1 | 0.025 | 0.01 | #DIV/0! | 2 | 0.025 | 0.01 | #DIV/0!
Percent Difference |  |  |  | #DIV/0! |  | Percent Difference |  |  |  | #DIV/0! | 1 | 0.01 | 0.2 | #DIV/0! | 2 | 0.01 | 0.2 | #DIV/0!
Upper 95% Confidence Interval for the Difference |  |  |  | #DIV/0! |  | Upper 95% Confidence Interval for the Difference |  |  |  | #DIV/0! | 1 | 0.01 | 0.15 | #DIV/0! | 2 | 0.01 | 0.15 | #DIV/0!
Lower 95% Confidence Interval for the Difference |  |  |  | #DIV/0! |  | Lower 95% Confidence Interval for the Difference |  |  |  | #DIV/0! | 1 | 0.01 | 0.1 | #DIV/0! | 2 | 0.01 | 0.1 | #DIV/0!
[11 empty] | 1 | 0.01 | 0.05 | #DIV/0! | 2 | 0.01 | 0.05 | #DIV/0!
Conclusions: Low Pool | [10 empty] | 1 | 0.01 | 0.025 | #DIV/0! | 2 | 0.01 | 0.025 | #DIV/0!
#DIV/0! | [10 empty] | 1 | 0.01 | 0.01 | #DIV/0! | 2 | 0.01 | 0.01 | #DIV/0!
#DIV/0!
#DIV/0!
Low Pool
Conclusions: High Pool | [11 empty] | alpha | 0.05
#DIV/0! | [11 empty] | DF | -2 |  | *=B34+D34-2
#DIV/0! | [11 empty] | t-value | #NUM! |  | *=T.INV.2T(O51,52)
#DIV/0! | [12 empty] | #DIV/0! |  | *=SQRT(((B36^2)/B34)+((D36^2)/D34))
[12 empty] | 95% CI | #NUM! |  | *=O53*O54
APPROVED: |  |  |  |  | Date: | [6 empty] | High Pool
[12 empty] | alpha | 0.05
[12 empty] | DF | -2 |  | *=H34+J34-2
Name: | [11 empty] | t-value | #NUM! |  | *=T.INV.2T(O58,O59)
[13 empty] | #DIV/0! |  | *=SQRT(((H36^2)/H34)+((J36^2)/J34))
Notes: | [11 empty] | 95% CI | #NUM! |  | *=O60*O61
1 | If there are missing data points (less than required replicates) repeat the experiment.
2 | Review data for outliers and exclude data with an assignable cause. If there is more than one outlier consider repeating the experiment.
3 | Conclusion based on statistical significance is provided for informational purposes. CLSI EP7 recommends interpretation based on clinical utility.
INTERFERENCE SCREENING
EXAMPLE

## Dose-Response Worksheet 1
| FIVE LEVEL INTERFERENCE DOSE RESPONSE STUDY |  |  |  | *Enter Dmax as absolute concentration, percent difference or both
ANALYTE 1
| Date: |  |  |  | Test Substance (Interferent)
| Analyte |  |  |  | Technician
| Units |  |  |  | *Maximum Allowable Difference (Dmax), Concentration |  |  |  |  | 0
| Method |  |  |  | *Maximum Allowable Difference (Dmax), Percent |  |  |  |  | %
Reagent Lot#
Calibrator Lot#
Instrument
Enter concentrations of Levels 1 and 5, and intermIxed levels will be calculated; Otherwise, enter concentrations for all five levels.
Interferent | Interferent | Rep 1 | Rep 2 | Rep 3 | Rep 4 | Rep 5 | Mean | Absolute | Absolute | Percent | Percent
Level | Concentration | Concentration | Concentration | Concentration | Concentration | Concentration | Concentration | Difference | Limit | Difference | Limit
1 | [8 empty] | Exceeded? |  | Exceeded?
2 | 0
3 | 0
4 | 0
5
Mean X | 0
Mean Y | #DIV/0! | [7 empty] | Xi-Xm^2 |  | Yi-Ym^2
Intercept | #DIV/0! | [6 empty] | X1 | 0 | Y1 | #VALUE!
Slope | #DIV/0! | [6 empty] | X2 | 0 | Y2 | #VALUE!
r | #DIV/0! | [6 empty] | X3 | 0 | Y3 | #VALUE!
r2 | #DIV/0! | [6 empty] | X4 | 0 | Y4 | #VALUE!
Sy/x | #DIV/0! | [6 empty] | X5 | 0 | Y5 | #VALUE!
Sum (Xi-Xm)^2 | 0
Sum (Yi-Ym)^2 | #VALUE! | [7 empty] | Control Line | (+) Dmax | (-) Dmax
SE(slope) | #VALUE! | [6 empty] | X1 | 0 | 0 | 0
N | 0 | [6 empty] | X5 | 0 | 0 | 0
t-statistic | #DIV/0! | [6 empty] | Y1 |  | #VALUE! | #VALUE!
DF | -2 | [6 empty] | Y5 |  | #VALUE! | #VALUE!
t-Comparison | #NUM!
Slope p≤0.05? | #DIV/0! | [7 empty] | Control Line | (+)%Dmax | (-)%Dmax
[8 empty] | X1 | 0 | 0 | 0
Conclusions: Dose-Response Experiment | [7 empty] | X5 | 0 | 0 | 0
#DIV/0! | [7 empty] | Y1 |  | #VALUE! | #VALUE!
[8 empty] | Y5 |  | #VALUE! | #VALUE!
Bias with Interferent Level 2 < the acceptable limit (Dmax), indicating that Interference is NOT Clinically Significant.
Bias with Interferent Level 3 < the acceptable limit (Dmax), indicating that Interference is NOT Clinically Significant.
Bias with Interferent Level 4 < the acceptable limit (Dmax), indicating that Interference is NOT Clinically Significant.
Bias with Interferent Level 5 > the acceptable limit (Dmax), indicating that Interference IS Clinically Significant.
POINT-TO-POINT INTERFERENCE CUTPOINT ESTIMATE | [8 empty] | Greater
Absolute Difference | Slope | Intercept | Interference |  | Percent Difference | Slope | Intercept | Interference | Abs or %
Level 1 to Level 2 |  |  |  |  | Level 1 to Level 2 |  |  |  | 0
Level 2 to Level 3 |  |  |  |  | Level 2 to Level 3 |  |  |  | 0
Level 3 to Level 4 |  |  |  |  | Level 3 to Level 4 |  |  |  | 0
Level 4 to Level 5 |  |  |  |  | Level 4 to Level 5 |  |  |  | 0
INTERPOLATED: |  | Level 1-2 | Level 2-3 | Level 3-4 | Level 4-5
Allowable Interferent Concentration |  | 0 | 0 | 0 | 0
APPROVED:
Name: |  |  |  |  | Date:
Notes:
1 | If there are more than two (2) missing data points (out of 25) consider repeating the experiment.
Alternatively, generate additional measurements to replace missing values; understanding that total measurement variability may increase.
2 | Review data for outliers and exclude data with an assignable cause. If there is more than one outlier consider repeating the experiment.
3 | Conclusion based on statistical significance is provided for informational purposes. CLSI EP7 recommends interpretation based on clinical utility.

## Dose-Response Worksheet 2
| FIVE LEVEL INTERFERENCE DOSE RESPONSE STUDY |  |  |  | *Enter Dmax as absolute concentration, percent difference or both
ANALYTE 2
| Date: |  |  |  | Test Substance (Interferent)
| Analyte |  |  |  | Technician
| Units |  |  |  | *Maximum Allowable Difference (Dmax), Concentration |  |  |  |  | 0
| Method |  |  |  | *Maximum Allowable Difference (Dmax), Percent |  |  |  |  | %
Reagent Lot#
Calibrator Lot#
Instrument
Enter concentrations of Levels 1 and 5, and intermIxed levels will be calculated; Otherwise, enter concentrations for all five levels.
Interferent | Interferent | Rep 1 | Rep 2 | Rep 3 | Rep 4 | Rep 5 | Mean | Absolute | Absolute | Percent | Percent
Level | Concentration | Concentration | Concentration | Concentration | Concentration | Concentration | Concentration | Difference | Limit | Difference | Limit
1 | [8 empty] | Exceeded? |  | Exceeded?
2 | 0
3 | 0
4 | 0
5
Mean X | 0
Mean Y | #DIV/0! | [7 empty] | Xi-Xm^2 |  | Yi-Ym^2
Intercept | #DIV/0! | [6 empty] | X1 | 0 | Y1 | #VALUE!
Slope | #DIV/0! | [6 empty] | X2 | 0 | Y2 | #VALUE!
r | #DIV/0! | [6 empty] | X3 | 0 | Y3 | #VALUE!
r2 | #DIV/0! | [6 empty] | X4 | 0 | Y4 | #VALUE!
Sy/x | #DIV/0! | [6 empty] | X5 | 0 | Y5 | #VALUE!
Sum (Xi-Xm)^2 | 0
Sum (Yi-Ym)^2 | #VALUE! | [7 empty] | Control Line | (+) Dmax | (-) Dmax
SE(slope) | #VALUE! | [6 empty] | X1 | 0 | 0 | 0
N | 0 | [6 empty] | X5 | 0 | 0 | 0
t-statistic | #DIV/0! | [6 empty] | Y1 |  | #VALUE! | #VALUE!
DF | -2 | [6 empty] | Y5 |  | #VALUE! | #VALUE!
t-Comparison | #NUM!
Slope p≤0.05? | #DIV/0! | [7 empty] | Control Line | (+)%Dmax | (-)%Dmax
[8 empty] | X1 | 0 | 0 | 0
Conclusions: Dose-Response Experiment | [7 empty] | X5 | 0 | 0 | 0
#DIV/0! | [7 empty] | Y1 |  | #VALUE! | #VALUE!
[8 empty] | Y5 |  | #VALUE! | #VALUE!
Bias with Interferent Level 2 < the acceptable limit (Dmax), indicating that Interference is NOT Clinically Significant.
Bias with Interferent Level 3 < the acceptable limit (Dmax), indicating that Interference is NOT Clinically Significant.
Bias with Interferent Level 4 < the acceptable limit (Dmax), indicating that Interference is NOT Clinically Significant.
Bias with Interferent Level 5 > the acceptable limit (Dmax), indicating that Interference IS Clinically Significant.
POINT-TO-POINT INTERFERENCE CUTPOINT ESTIMATE | [8 empty] | Greater
Absolute Difference | Slope | Intercept | Interference |  | Percent Difference | Slope | Intercept | Interference | Abs or %
Level 1 to Level 2 |  |  |  |  | Level 1 to Level 2 |  |  |  | 0
Level 2 to Level 3 |  |  |  |  | Level 2 to Level 3 |  |  |  | 0
Level 3 to Level 4 |  |  |  |  | Level 3 to Level 4 |  |  |  | 0
Level 4 to Level 5 |  |  |  |  | Level 4 to Level 5 |  |  |  | 0
INTERPOLATED: |  | Level 1-2 | Level 2-3 | Level 3-4 | Level 4-5
Allowable Interferent Concentration |  | 0 | 0 | 0 | 0
APPROVED:
Name: |  |  |  |  | Date:
Notes:
1 | If there are more than two (2) missing data points (out of 25) consider repeating the experiment.
Alternatively, generate additional measurements to replace missing values; understanding that total measurement variability may increase.
2 | Review data for outliers and exclude data with an assignable cause. If there is more than one outlier consider repeating the experiment.
3 | Conclusion based on statistical significance is provided for informational purposes. CLSI EP7 recommends interpretation based on clinical utility.

## Dose-Response Worksheet 3
| FIVE LEVEL INTERFERENCE DOSE RESPONSE STUDY |  |  |  | *Enter Dmax as absolute concentration, percent difference or both
ANALYTE 3
| Date: |  |  |  | Test Substance (Interferent)
| Analyte |  |  |  | Technician
| Units |  |  |  | *Maximum Allowable Difference (Dmax), Concentration |  |  |  |  | 0
| Method |  |  |  | *Maximum Allowable Difference (Dmax), Percent |  |  |  |  | %
Reagent Lot#
Calibrator Lot#
Instrument
Enter concentrations of Levels 1 and 5, and intermIxed levels will be calculated; Otherwise, enter concentrations for all five levels.
Interferent | Interferent | Rep 1 | Rep 2 | Rep 3 | Rep 4 | Rep 5 | Mean | Absolute | Absolute | Percent | Percent
Level | Concentration | Concentration | Concentration | Concentration | Concentration | Concentration | Concentration | Difference | Limit | Difference | Limit
1 | [7 empty] | #VALUE! | Exceeded? | #VALUE! | Exceeded?
2 | 0 | [6 empty] | #VALUE! |  | #VALUE!
3 | 0 | [6 empty] | #VALUE! |  | #VALUE!
4 | 0 | [6 empty] | #VALUE! |  | #VALUE!
5 | [7 empty] | #VALUE! |  | #VALUE!
Mean X | 0
Mean Y | #DIV/0! | [7 empty] | Xi-Xm^2 |  | Yi-Ym^2
Intercept | #DIV/0! | [6 empty] | X1 | 0 | Y1 | #VALUE!
Slope | #DIV/0! | [6 empty] | X2 | 0 | Y2 | #VALUE!
r | #DIV/0! | [6 empty] | X3 | 0 | Y3 | #VALUE!
r2 | #DIV/0! | [6 empty] | X4 | 0 | Y4 | #VALUE!
Sy/x | #DIV/0! | [6 empty] | X5 | 0 | Y5 | #VALUE!
Sum (Xi-Xm)^2 | 0
Sum (Yi-Ym)^2 | #VALUE! | [7 empty] | Control Line | (+) Dmax | (-) Dmax
SE(slope) | #VALUE! | [6 empty] | X1 | 0 | 0 | 0
N | 0 | [6 empty] | X5 | 0 | 0 | 0
t-statistic | #DIV/0! | [6 empty] | Y1 |  | #VALUE! | #VALUE!
DF | -2 | [6 empty] | Y5 |  | #VALUE! | #VALUE!
t-Comparison | #NUM!
Slope p≤0.05? | #DIV/0! | [7 empty] | Control Line | (+)%Dmax | (-)%Dmax
[8 empty] | X1 | 0 | 0 | 0
Conclusions: Dose-Response Experiment | [7 empty] | X5 | 0 | 0 | 0
#DIV/0! | [7 empty] | Y1 |  | #VALUE! | #VALUE!
[8 empty] | Y5 |  | #VALUE! | #VALUE!
Bias with Interferent Level 2 < the acceptable limit (Dmax), indicating that Interference is NOT Clinically Significant.
Bias with Interferent Level 3 < the acceptable limit (Dmax), indicating that Interference is NOT Clinically Significant.
Bias with Interferent Level 4 < the acceptable limit (Dmax), indicating that Interference is NOT Clinically Significant.
Bias with Interferent Level 5 > the acceptable limit (Dmax), indicating that Interference IS Clinically Significant.
POINT-TO-POINT INTERFERENCE CUTPOINT ESTIMATE | [8 empty] | Greater
Absolute Difference | Slope | Intercept | Interference |  | Percent Difference | Slope | Intercept | Interference | Abs or %
Level 1 to Level 2 |  |  |  |  | Level 1 to Level 2 |  |  |  | 0
Level 2 to Level 3 |  |  |  |  | Level 2 to Level 3 |  |  |  | 0
Level 3 to Level 4 |  |  |  |  | Level 3 to Level 4 |  |  |  | 0
Level 4 to Level 5 |  |  |  |  | Level 4 to Level 5 |  |  |  | 0
INTERPOLATED: |  | Level 1-2 | Level 2-3 | Level 3-4 | Level 4-5
Allowable Interferent Concentration |  | 0 | 0 | 0 | 0
APPROVED:
Name: |  |  |  |  | Date:
Notes:
1 | If there are more than two (2) missing data points (out of 25) consider repeating the experiment.
Alternatively, generate additional measurements to replace missing values; understanding that total measurement variability may increase.
2 | Review data for outliers and exclude data with an assignable cause. If there is more than one outlier consider repeating the experiment.
3 | Conclusion based on statistical significance is provided for informational purposes. CLSI EP7 recommends interpretation based on clinical utility.

## Calculations (hidden)
DOSE-RESPONSE CALCULATIONS FIVE POINTS: ANALYTE 1 | [7 empty] | DOSE-RESPONSE CALCULATIONS FIVE POINTS: ANALYTE 2 | [7 empty] | DOSE-RESPONSE CALCULATIONS FIVE POINTS: ANALYTE 3
Interferent | Concentration | (Xi-Xm)^2 | [5 empty] | Interferent | Concentration | (Xi-Xm)^2 | [5 empty] | Interferent | Concentration | (Xi-Xm)^2
Concentration |  |  | Pred Conc | 95%CI | Low CL | High CL | (Yi-Yp)^2 | Concentration |  |  | Pred Conc | 95%CI | Low CL | High CL | (Yi-Yp)^2 | Concentration |  |  | Pred Conc | 95%CI | Low CL | High CL | (Yi-Yp)^2
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!

## Sheet1 (hidden)
DOSE-RESPONSE CALCULATIONS FIVE POINTS: ANALYTE 1
| Interferent | Concentration | (Xi-Xm)^2
| Level |  |  | Pred Conc | 95%CI | Low CL | High CL | (Yi-Yp)^2 | (Yi-Ym)^2 | (Yp-Ym)^2
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
#VALUE! |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
Mean | #DIV/0! | #DIV/0!
SD | #DIV/0! | #DIV/0!
Sum | 0 | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #VALUE! | #REF!
|  | SSTotal= | #VALUE!
|  | SSResidual= | #REF!
|  | SSRegression= | #VALUE!
|  | r^2= | #VALUE!
|  | r= | #VALUE!
|  | slope= | #REF!
|  | intercept= | #DIV/0!
|  | Sy/x= | #REF!